## Slide 1: Rapid Antibody Based Blood Test for COVID-19
Learning Resource Package for Lab Technicians

## Slide 2
Why do we need Antibody tests?
Antigens & antibodies are considerably more stable than viral RNA
Less susceptible to spoliation during transport and storage, therefore reducing the chance of false-negative results
Testing accuracy is also improved as antigens and antibodies are more uniformly available in sputum and blood samples
Ability to detect past infections as virus- specific antibodies (unlike viral RNA) can persist in the bloodstream for many years after infection and thus help in RETROSPECTIVE DETECTION
Type of antibody and its relative levels could also be used to indicate the stage of infection and estimate time since exposure for contact tracing

## Slide 3: Role of rapid antibody based blood tests in COVID-19
Supplementary tool to assess the prevalence of the diseases within a specific area/parameter
Only be of utility after a minimum of seven days of onset of symptoms
Data about these tests is emerging and understanding of their utility for diagnosis is still evolving
Useful for epidemiological studies and surveillance purposes
To re-emphasize, antibody detection is NOT used for DIAGNOSIS as per current guidelines

## Slide 4: Strategy for COVID19 testing in India (Version 4, dated 09/04/2020)
  All symptomatic individuals who have undertaken travel in affected state the last 14 days
  All symptomatic contacts of laboratory confirmed cases
  All symptomatic health care workers
  All patients with Severe Acute Respiratory Illness (fever AND cough and/or shortness of breath)
  Asymptomatic direct and high-risk contacts of a confirmed case should be tested once between day 5 and day 14 of coming in his/her contact
  In hotspots/cluster (as per MOHFW) and in large migration gatherings/ evacuees centers
  All symptomatic Influenza Like Illness (ILI) (fever, cough, sore throat, runny nose)
    Within 7 days of illness – RT-PCR
    After 7 days of illness – Rapid Antibody test

## Slide 5
  Strategy for use of rapid antibody based blood test

## Slide 6: Serological Surveillance of COVID-19 using Antibody Tests-Validation guidelines
Important to Validate the accuracy of kits
Test is a complimentary tool to actual diagnosis of the infection/illness using the approved technique of RT-PCR
Use these kits in patients who have already been confirmed to have infection using the RT-PCR and are in various stages of illness
Persons who can be considered to be at high risk of contracting/ having contracted the infection from these confirmed positive patients
Target population can be classified into five categories so as to aid the interpretation of the results of the antibody test
It is important to correlate serological findings with various stages of the infection/ illness
* Source: Health & Family Welfare Department , Government of Kerala – COVID-19 Guidelines for conducting tests to identify community transmission

## Slide 7: Serological Surveillance of COVID-19 using Antibody Tests-Validation guidelines
1
PCR-positive Cured / Recovered Patients
2
PCR-positive Extremely Ill Patients in ICU
3
4
PCR-positive Moderately Ill Patient
PCR-positive Mildly Ill Patients (Early diagnosed)
5
PCR-negative / undiagnosed Close Primary Contacts of
Persons with first contact
less than 7 days ago
between 7 and 14 days
more than 14 days ago

## Slide 8
Rapid IgM-IgG Combined Antibody test for COVID-19
  Rapid IgM-IgG Combined Antibody Test for COVID-19 is a visually read & used to qualitatively detect IgG and IgM antibodies
  Human finger-stick (capillary) or venous whole blood, serum, and plasma samples are used
  Developed for COVID -19 testing to contain the spread of infection
  To provide reliable diagnosis to all individuals meeting the inclusion criteria of COVID-19 testing

## Slide 9
Test Principle
Test card contains:
Novel coronavirus antigen and quality control antibody colloidal gold Marker
Two detection lines (G and M lines) and one quality control line (C)
M is fixed with monoclonal anti-human IgM antibody for detecting the novel coronavirus IgM antibody
G is fixed with monoclonal antihuman IgG antibody for detecting the novel coronavirus IgG antibody
The quality control antibody is fixed on the C line

## Slide 10
Test Principle
Test card contains:
Colloidal gold-labeled recombinant novel coronavirus antigen and quality control antibody colloidal gold Marker
Two detection lines (G and M lines) and one quality control line (C) fixed on a nitrocellulose membrane
M is fixed with monoclonal anti-human IgM antibody for detecting the novel coronavirus IgM antibody
G is fixed with monoclonal antihuman IgG antibody for detecting the novel coronavirus IgG antibody
The quality control antibody is fixed on the C line

## Slide 11
Test Principle
After adding appropriate amount, the test sample will move forward along the test card via capillary action
If the sample contains IgM antibody, antibody will bind to the colloidal gold labeled novel coronavirus antigen
Antibody/antigen complex will be captured by the anti-human IgM antibody immobilized on the membrane, forming a red M line and indicating a positive result for IgM antibody
Similarly, If the sample contains IgG antibodies, a red G line will be formed indicating a positive result for IgG antibody

## Slide 12
Test Principle – contd.
If neither antibody is present, a negative result is displayed. The card also contains a quality control line (C)
Regardless of what antibodies are present the C line should appear to indicate that the sample has been transported properly through the membrane
If the C line does not appear it indicates that the test result is invalid and a new, unopened test cassette is required to retest the sample

## Slide 13: Suggested Testing Window
Variation of the Levels of SARS-CoV-2 RNA and Antigen, IgM and IgG after infection

## Slide 14: COVID-19 IgM-IgG Rapid Test Kit
Test cassette
Sample diluent
Dropper
Lancet device
Alcohol Pad

## Slide 15
4 Simple Steps

## Slide 16
Step 1: Using capillary sampler, obtain 20μL of fingerstick or venous whole blood specimen or 10μL of serum or plasma
Step 2: Dispense the specimen into the Test Cassette sample well. Ensure that the entire sample is dispensed into the sample well
Step 3: Remove cap of the Buffer Solution bottle and dispense 2-3 drops into the Test Cassette sample well
Step 4: Allow test to run for 10 minutes. Read the results by viewing the detection window. Test results that have run over 15 minutes are invalid
4 Simple Steps

## Slide 17: Interpretation of Results
A total of three detection lines are possible, the control with (C) line appearing when sample has finished flowing through the cassette, indicating the test is complete
In next four slides, each test reading is interpreted separately
C line will always appear if the sample has been transported properly through the membrane

## Slide 18: Interpretation of Results
Negative Result: If only the quality control line (C) appears and the detection lines G and M are not visible, then no novel coronavirus antibody has been detected and the result is negative
Indicates that the person has not been recently exposed to the COVID-19 virus

## Slide 19: Interpretation of Results
IgG only positive: If both the quality control line (C) and the detection line G appears, then the novel coronavirus IgG antibody has been detected and the result is positive for the IgG antibody
Indicates that either infection is over in the person or is in its last stage

## Slide 20: Interpretation of Results
IgM only positive: If both the quality control line (C) and the detection line M appears, then the novel coronavirus IgM antibody has been detected and the result is positive for the IgM antibody
Indicates person is likely in early stage of infection

## Slide 21: Interpretation of Results
IgG and IgM positive: If the quality control line (C) and both detection lines G and M appear, then the novel coronavirus IgG and IgM antibodies have been detected and the result is positive for both IgG and IgM antibodies
Indicates that the person is likely in the middle stage of infection

## Slide 22: Warnings and Precautions
Only for human in vitro clinical diagnostics only
After opening the sealed cassette pouch, test should be done within one hour
Do not immerse test cassette in water
Do not freeze test cassette or buffer solution
Wear protective PPE (Gloves & Mask)
Wash hands thoroughly after handling specimens
Do not use test cassette, buffer solution, or any kit component beyond the indicated expiration date
Dispose of all used or damaged test cassettes, capillary samplers, or other kit component as biohazardous materials (Red colored dustbin)
Do not use test cassette, buffer solution, or any other kit components if the pouch is damaged or the seal is broken
Do not use samples containing lipids, hemolysis, or turbidity which can affect results

## Slide 23: Other Instructions
Storage Instructions
Reagent should be stored in the dark at room temperature (2° to 45°C) and has a shelf-life of 12 months.
Container should be protected from light after being opened. Do not freeze.
Sample Requirements
Suitable for human serum, plasma, or whole blood samples (capillary or venous) including samples prepared by commonly used anticoagulants (EDTA, heparin, sodium citrate).
Fresh samples should be collected and tested immediately.
Serum and plasma samples can be stored at 2-8°C for 5 days.
Anticoagulated whole blood samples can be stored at 2-8°C for 7 days.
Before testing, samples stored in refrigerated or frozen storage should be slowly returned to room temperature (15-30° C) and stirred.
When particulates are clearly visible in the sample the precipitate should be removed by centrifugation before testing.

## Slide 24
Aren’t suitable for diagnosing active COVID-19 infections
Body’s antibody response to COVID-19 is slow – considerably slower than we might expect
While data at this point is still limited, initial IgM antibody response doesn’t peak until ~9 days after initial infection and the IgG antibody response doesn’t peak until day ~12
Most viruses take 5 days to produce initial immune response
Rapid test can ensure the immunity of people but can not provide the information that they are no longer spreading the virus. For that purpose, RT-PCR plays an important role
May give false negative results – if done in window period before the antibody appears (antibody takes 8-14 days to develop following natural infection)
  Limitations and Challenges

## Slide 25: Limitations and Challenges- contd.
Technical challenges in making antigens: Difficult to make the protein in laboratories using cell cultures which forms the viral coat, essential to trigger the immune response to produce antibodies for inclusion in immunoassays. There are questions which antigens [proteins] are the best to produce antibodies
Nucleocapsid protein: the most abundant viral protein, which means it’s easy to detect, but less specific
Spike protein: more specific, more divergent but which part of it to be used is less obvious
False positive results: due to cross reactivity to other coronavirus infections; amino acids of the spike proteins of SARS and the COVID-19 virus, there’s a 75% probability and overall figure for common cold-causing coronaviruses is about 50–60%.

## Slide 26: State specific Action Points
Rapid antibody tests approved by US-FDA/ CE-IVD or non-CE-IVD validated by ICMR-NIV with marketing approval by DCGI to be used
In order to ensure that all such cases are monitored and necessary action is initiated with respect to infectious disease management, details of all test results shall be uploaded in ICMR portal
All such organizations are duty bound to register themselves on ICMR portal and upload the data in real-time
In case of failure to do so, they will be held liable to action under Disaster Management Act, 2005
Follow up action to be determined as per the State protocols
Govt. of Kerala COVID-19 guidelines for conducting tests to identify community transmission can used as reference

## Slide 27: References
Indian Council of Medical Research https://icmr.nic.in/content/covid-19
Rapid IgM-IgG Combined Antibody test for Coronavirus brochure, Biomedomics. (Approved by ICMR) https://www.oxfordbiosystems.com/Portals/0/PDF/Biomedomics/COVID19_flyer.pdf
COVID-19 Guidelines for conducting tests to identify community transmission, Health & Family Welfare Department , Government of Kerala